NRD doesn’t apply the same filters/manipulations that you see on .NCS files

Files needed:

- RawData.nrd

- Config file (p98cs_replay_sc.cfg) <-- will need to update file paths (top of script)

  - MICROS - SetSubSamplingInterleave should always be 1 (we don’t subsample from 32k on micro channels)

  - MACROS – The SubSampling is typically at 16 (corresponds to 2k Hz)

  - Replay_Atlas_generic.cfg -- use this to make config file

    - make separate file for the .nrd and .cfg file within nrd_replay folder

    - change SetDataDirectory “C:\dataReplay\XXXX” to the new folder location and RawDataFile “XXX-SPECIFY.nrd” to the specific file location

      - copy address as text

    - delete extraneous stuff between “-SetNetComDataBufferSize Events 3000” and “-CreateCscAcqEnt”

    - copy paste the “-CreateCscAcqEnt” part to correspond w/ # of CSC channels

      - 0-255 or 1-256 are all macros for UC Davis data

    - delete overlap between micros/macros -> save file

      - delete macros if necessary?

    - MAKE SURE YOU UPLOAD THE CONFIG FILE WITHIN FOLDER YOU JUST MADE (CHECK UPDATED TIME)

Software needed:

- Also need Pegasus software (skip license file, we’re not using this for actual recording, just getting data)

  - open + start with new data file/setup

- MATLAB Script – NetCome_expStarter_ATLAS_AC.m (that interfaces with Pegasus software)

  - Calls NLX helper files from package folder “neuralynxNetcom201” (under nrd replay, make sure to add this to the path)

  - If you redownload NetCome_expStarter_ATLAS_AC.m (Nov 26, 2024), make sure to comment out lines 17 + 24, and update the timestamps in the code as usual

After uploading config file to Pegasus

- Pegasus software > Raw Data file Playback Properties

  - Filename – update this to what you did in the config file

  - UNCHECK continuous file playback

  - Change playback speed to the SLOWEST (otherwise we might have dropped samples)

  - Timestamp range – update matlab script (line 23) with what you see in pegasus software

- Pegasus Software > Add Window > Time Window > Display > Add Plots > Add like 15 CSC rows

  - Micros: View > Acquisition Entities and Display Properties > Under entities > CSC (#s should align with config file) > Select all the CSC (micros) > Sample freq should be 32k Hz, Input range should be 3000 uV, sub-sampling is 1, Low cut is 0.1 Hz, No High cut applied

  - Macros: Input rage should be 10k uV, Remove high Cut as well, sub sampling interleave 16

- Run MATLAB Script > Then on Pegasus, ACQ turns green, REC turns purple

  - If both are gray, just click the REC button so that both are green/purple respectively

  - make sure you drag expStarter file into new folder

When finished replaying

A folder will be generated with today’s date/time

- Folder has a bunch of NCS files

- Also has event file (for TTLs)